## Slide 1: TEM, HREM & Thoughts
Image Curtesy Staf Van Tenderloo

## Slide 2: Why HREM
In order to…
  Make your boss/advisor happy
  Obtain a pretty picture for publication
  Obtain local structural information
  Obtain precise structural information atomic column by atomic column
  Obtain some new science (and not reinvent)
  Follow the science, not the electron

## Slide 3: Two fundamental modes
Low Resolution interpretation
Local periodicities
Defects such as twins etc
High Resolution interpretation
Local positions at defects etc with 0.1 Angstrom accuracy or better

## Slide 4: Examples
MTP (1983)
SrTiO3 (001) (2015)
1nm
TiO2 DL

## Slide 6: All Data

## Slide 7
Low Resolution …

## Slide 8
Higher resolution …

## Slide 9
Higher resolution …

## Slide 10
Higher resolution …

## Slide 11
Higher resolution …

## Slide 12: All Data

## Slide 13: But…
Classic resolution limit is approximately the radiation wavelength
Electron wavelength is ~ 0.025 Angstroms
Best resolution to date is ~ 0.5 Angstroms, 20 time wavelength
Lens Aberrations

## Slide 14: Aberrations
Microscopes without correctors have lenses about as good as a milk bottle
  Distortions
  Some spacings show but should not
  Some spacings should show but do not
  Information is often not where it should be

## Slide 15: Defocus
Sample
Focal Plane
Different distances = different phase
Obj.
lens
Gaussian
focus

## Slide 16: Spherical Aberration Cs
Obj.
lens
Sample
Gaussian
focus
least
confusion
Focal Plane

## Slide 17: Chromatic Aberration
Disc of Least Confusion
Focal length of lens is dependent upon wavelength

## Slide 18: Aberrations
Aberrations are a phase-shift in reciprocal space
Multiply by exp(-ic(u))
Can expand as a Taylor series
c(u) = A + Bu + Cu2 + D|u.a|2 + …
A,B don’t matter, C is defocus, D is astigmatism, Cs has u4
Image
Plane
Optic axis
c(u)
u
Phase distance from surface of constant phase

## Slide 19: Aberrations
C1,0
C1,2
C2,1
C2,3
C4,1
C4,3
C4,5
C3,0
C3,2
C3,4
C5,0
C5,2
C5,4
C5,6
Surfaces of constant phase

## Slide 20: More complete formalism
. If y(r) is the wave exiting the specimen, the imaging
process is
Fourier Phase Fourier
Transform Change Transform
y(r) --------------> Y(u) -----------> Y'(u) --------------> y'(r) -
Incoherent
Intensity Average
---------> |y'(r)|2 ---------> I(r), Final Image.

## Slide 21: Linear & Non-Linear

## Slide 22: Linear Contrast Model
Non-linear: General solution
I(r) = ò exp(2piv.r) P(v)dv
P(v) = <ò Y*(u-v) Y(u) exp(-i c(u) + i c(u-v))du>
= ò Y*(u-v) Y(u) exp(-i c(u) + i c(u-v))E(u,u-v)du
E(u,u-v) = exp(-|Ñc(u)-Ñc(u-v)|2/4a)exp(-p2l2(u2-|u-v|2)2/4ß)
Involves both a statistical average “<..>” and an Envelope Term,
typically a first-order approximation.
Linear: Thin sample
I(r) = ò cos(2pv.r) E(v,0)sin(c(v)) dv

## Slide 23: Astigmatism & Drift
Drift
Well aligned
Misaligned

## Slide 24
Effect of slight
beam tilt

## Slide 25: Coma-Free Alignment
Before
After
+a
-a
+a
-a

## Slide 26: Beam Tilt
Diffractograms
(FFT patterns of amorphous regions)
Similar to the effect of astigmatism
Astigmatism correction may be required after the alignment of beam tilt.

## Slide 27: Translation
Obj.
lens
Sample
Gaussian
focus
least
confusion
Focal Plane

## Slide 28
C
Ghost Image

## Slide 30: Au@NdScO3
Unfriendly sample – large differences in depth of nanoparticles leads to different contrast

## Slide 31: Direct Interpretation
Image Curtesy Van Tenderloo

## Slide 32: Direct Interpretation
Strain Contrast!
Image Curtesy Van Tenderloo

## Slide 33: Caveats
Linear imaging is an approximation that is a guide only.
Kinematical is an approximation that is rarely good enough
Pseudo-Linear Imaging
Pseudo-Kinematical
Rigor Please!

## Slide 34: Examples

## Slide 35: Interfaces
Columnar grain boundaries in SrTiO3 film are associated with {111} planar defects in the SrRuO3 layer. Two atomic structural models of the antiphase boundaries in the (-110) plane, with a crystallographic shear vector R=a/2[001] (a) and R=a/2[-1-11] (b).
a/3[-112]  a/6[-221] + a/2[001]

## Slide 38
Si
a-BN
h-BN
30 Å

## Slide 41: Structural Fluctuations (Iijima)
P. M. Ajayan, L. D. Marks,. 24-6, 229 (1990)

## Slide 42
High-Resolution TEM: <1 nm
Uniformity in size (<1nm) and spatial distribution
High crystallinity evident from visible lattice fringes inside particles
Lattice spacing = 2.39  0.07(Å)
Icosahedral shape based on trace analysis of particle edge
HREM image of Au13(PPh3)4(SC12)4
Huiping Xu, Ray Twesten

## Slide 44: Nanoparticle surfaces?
Oleic Acid Acetic Acid
1nm
PRL 111, 156101 (2013)
1nm

## Slide 45: SrTiO3 (001) surface
[100]
[110]
thickness increase
1nm
x
y
[001]
[110]
[110]
Around 20 nm size
Surface is flat in projection

## Slide 46: Profile Imaging: Fit Bulk
Area 2
Area 1
ACAT at ANL Bulk Fit

## Slide 47: Different Surfaces
DFT positions (PBEsol) used in multislice
Can use DFT potentials as well
MacTempas

## Slide 48: Nanoparticle surfaces
Oleic Acid Acetic Acid
1nm
SrO surface
PRL 111, 156101 (2013)
1nm
TiO2 DL

## Slide 49: Nanoparticle surfaces?
1nm
SrO surface
Precision good enough – functional matters!
Chart: Interlayer numbers (i)
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
PBEsol0 | 1.7755 | 2.0132 | 1.939 | 1.9438 | 1.9794 | 1.9282 | 1.9905 | 1.9182 | 1.9913
EXP | 1.7529 | 2.0748 | 1.9533 | 1.924 | 2.0071 | 1.9094 | 1.9923 | 1.9165 | 1.9918
PBE | 1.742 | 2.042 | 1.9141 | 1.9504 | 1.9678 | 1.9283 | 1.9863 | 1.9181 | 1.9906
To Bulk
PRL 111, 156101 (2013)

## Slide 50: Microscopy Big Bang
Photon
Scanned photon
Electron
Scanned electron
Aberration –corrected electron
Other
Curtesy A. Howie

## Slide 51: Spatial Resolution

## Slide 52
The Big Bang
TEAM Project: R. Erni, et al., Phys Rev Lett, 102 (2009) 96101
43 meV 40kV & 0.12nm
(Bell, 2010)
Atomic imaging and EELS, R. Keile, UIC
3 meV
(Krivenak, 2021)
0.2nm SEM of a surface, J Ciston
3D atomic tomography of a dislocation C-C. Chen et al
Combined effect of light & H2O, P. Crozier

## Slide 53: Resolution Limiting Factors.
Incoherent Aberrations
( Source Size, Energy Spread,
HT stability, lens stability
Chromatic Aberration ).
Coherent Aberrations
(Objective lens
Spherical Aberration)
Detectors
(Phosphor coupled CCD)
(Monochromation,
Novel Sources, improved stabilities, Cc correction)
(Spherical Aberration Correctors)
(Direct Electron Detectors)

## Slide 54: Aberrations
C1,0
C1,2
C2,1
C2,3
C4,1
C4,3
C4,5
C3,0
C3,2
C3,4
C5,0
C5,2
C5,4
C5,6

## Slide 55: Basic concept
One cannot correct with a spherical lens, i.e. one with only even u terms
One can, however, generate –ve 4th order terms by breaking symmetry
  Use hexapoles, which introduce terms with 3-fold symmetry
  Octapoles, 4-fold symmetry
  Arrange to cancel out most of the 3-fold or 4-fold aberrations

## Slide 56: Multipole Lenses

## Slide 57
Correction of Spherical Aberration.
Image plane
Corrector
Objective
lens
Object
plane
Image plane
Need to have non-spherical elements

## Slide 58: Practical Approach
Convert beam from radially symmetric to ellipsoidal (one direction), compensate Cs for this direction, switch by 90 degrees (or 60) and repeat

## Slide 59
Corrector Optics.
Dr Peter Hartel, CEOS

## Slide 60: Chromatic Aberration
Hexapole Wien corrector: Balance Lorentz (magnetic) and Electrostatic Force for some energy
Electrons of lower (higher) energy have a net deflection

## Slide 61: Two early aberration correctors
Sextupole - round lens - sextupole corr. (~1980)
(grandfather to CEOS’s corrector)
Quadrupole-octupole corrector (~1975)
(grandfather to Nion’s corrector)
Neither corrector ever worked properly, most likely because parasitic aberrations were neither quantified nor accurately compensated.

## Slide 62: Aberrations
Corrects all aberrations up to C5,0;
first strong aberrations are 7th order ones
C1,0
C1,2
C2,1
C2,3
C4,1
C4,3
C4,5
C3,0
C3,2
C3,4
C5,0
C5,2
C5,4
C5,6
corrected
(or minimized)
aberrations

## Slide 63
Instrumentation – Mark 1 (JEOL 2200FS).

## Slide 64: Instrumentation – Version 2 (JEOL 2200 MCO).

## Slide 65
P.S
CESCOR
P.S
CETCOR
Linux
CESCOR
CETCOR
LAN
EDS PC
GATAN camera PS
FIBER OPTIC
STEM PC
JEOL ASID
TEM PC
JEOL TEMCON
TEM Server
JEOL Ext
Screen Monitor
HUB
Cross cable
DigiScan
Hamamatsu
Control
CESCOR
CETCOR
CCD
Video
EOS
LAN
LAN
LAN
LAN
LAN
IEEE1394
IEEE 1394
VME
HT tank
FIBER
OPTIC
flat cable
(LENS DEF)
ARC NET
CAMERA PC
CEOS Viewer
Gatan
JEOL Ext
Microscope handler
System Overview.

## Slide 66
ADL
TL21
TL22
TL11
TL12
Objective
HP 1
HP 2
DP11
DP12
QP
HP
DPHP1
DP21
DP22
DPHP2
IShift
Dshift
QP
A1, A2 correction
B2 correction,
Beam centering
Image and diffraction plane
alignment
Image shift
B2 correction,
Beam centering
Diffraction shift
A1 diffraction plane correction
Hexapole centering
Cs correction
2nd Order correction (off, on axis)
- Anisotropic coma

## Slide 67: NU ARM

## Slide 68
Sawada, H. et al. Experimental evaluation of a spherical aberration corrected TEM and STEM, Journal of Electron Microscopy, 2005, 52, 2, 120,

## Slide 69: Cs correction

## Slide 70: TEM of Graphene at 80 kV

## Slide 71: Aberration Correction in JEM-2200FS
1 nm
1 nm
Hexapole off
Hexapole on
Ronchigram
Simulated probe
Conventional
Ip = 10 pA
0.16 nm(HM)
0.76 nm(TM)
Cs-corrected
Ip = 50 pA
0.12 nm(HM)
0.35 nm(TM)
HAADF-STEM image
Si(110)
11 mrad
34 mrad
25 mrad

## Slide 72
Intensity profile
Histogram
FFT
444(78pm)
624(72pm)
804(60pm)
713(70pm)
Intensity profile from FFT spot
Si(112) 78 pm dumbbell image

## Slide 73: 30 kV
Experimental and calculated Cc/Cs-corrected 30-kV HRTEM images of graphene (A) and MoS2 (B). The diffractograms (A1 and B1) of the experimental raw images of graphene (A2) and MoS2 (B2) in bright-atom contrast (field of view 40×40 nm2) indicate the achieved resolution. The magnified images (A3 and B3) directly allow for identifying the atomic structure. The averaged experimental images (A4 and B4) show a strong signal-to-noise improvement and are in good agreement with the simulated images (A5 and B5) as visible in the corresponding line profiles (A6 and B6). The scale bars in A5 and B5 correspond to 0.5 nm.
Chromatic Aberration Correction for Atomic Resolution TEM Imaging from 20 to 80 kV
Linck et al, Phys. Rev. Lett. 117, 076101, 2016

## Slide 74: Beware!
All you get with advanced instruments is easier/quicker alignment and better signal/images
Many people sell the advantages of their instrument/technique (42, Douglas Adams)
It is easier for a novice to get a “good” image, but that does not make it correct or representative
There is no “dynamical diffraction” corrector
Nothing will correct for a bad sample
Nothing will correct for a bad interpretation